## REPORT SUMMARY
PARAMETER / SECTION | VERIFIED METROLOGICAL DATA
OVERALL VERIFICATION STATUS | PASSED (COMPLIANT)
Report Number | R76-2026-000073
Public Verification ID | R76-2026-D6F438AF
Verification Route URL | http://localhost:5173/public/report/R76-2026-D6F438AF
Issue Date | 2026-09-26
Regulatory Baseline | OIML R 76-1:2006 (TYPE_EVALUATION)
Compliance Determination | All 5 applicable OIML R 76-1:2006 metrological tests have been verified compliant with Maximum Permissible Errors.
Verification Laboratory | Central Legal Metrology Testing Laboratory
Laboratory Location | Directorate of Legal Metrology
Authorized Testing Officer | Pramod Patil
Instrument Manufacturer | Mettler Toledo
Model Number | ICS425
Serial Number | MT-55198
Instrument Type | Non-Automatic Weighing Instrument (Compact Industrial Scale)
Accuracy Class | Class III
Maximum Capacity (Max) | 15 kg
Minimum Capacity (Min) | 0.1 kg
Verification Scale Interval (e) | 0.005 kg
Actual Scale Interval (d) | 0.001 kg
Resolution (n = Max/e) | 3000
Engineering Unit | kg
Ambient Temperature | 21.8 deg C
Relative Humidity | 48 %
Atmospheric Pressure | 1013.25 hPa

## INSTRUMENT DETAILS
SPECIFICATION PARAMETER | NOMINAL / RECORDED VALUE | REGULATORY REQUIREMENT / CLAUSE
Manufacturer | Mettler Toledo | Clause 3.1 & 7.1 Identification
Model Designation | ICS425 | Clause 7.1 Descriptive Markings
Serial Number | MT-55198 | Unique Traceability Marking
Instrument Type | Non-Automatic Weighing Instrument (Compact Industrial Scale) | Clause 2.1 Non-Automatic Weighing Instrument
Accuracy Class | Class III | Clause 3.2 Accuracy Classification
Maximum Capacity (Max) | 15 kg | Clause 3.3.1 Max Specification
Minimum Capacity (Min) | 0.1 kg | Clause 3.3.2 Min Specification
Verification Scale Interval (e) | 0.005 kg | Clause 3.3.3 Verification Interval
Actual Scale Interval (d) | 0.001 kg | Clause 3.3.4 Scale Interval
Resolution (n = Max/e) | 3,000 intervals | Clause 3.2 Table 3 Interval Limits
Primary Engineering Unit | kg | Clause 3.4 Permissible Metric Units
Device Indication Type | Digital Electronic Indicator | Clause 4.3 Digital Indicating Device

## ENVIRONMENTAL CONDITIONS
ENVIRONMENTAL PARAMETER | OBSERVED TEST CONDITION | OIML R-76 PERMISSIBLE RANGE | STATUS
Ambient Temperature | 21.8 deg C | +10 deg C to +40 deg C (Class III standard) | PASS
Relative Humidity | 48 % | <= 85% Non-condensing | PASS
Atmospheric Pressure | 1013.25 hPa | 860 hPa to 1060 hPa (Standard barometric) | PASS
Thermal Stability | Equilibrated > 8 hours | Stable ambient condition during test | PASS

## WEIGHING PERFORMANCE
Test No | Phase | Applied Load L (kg) | Indication I (kg) | Changeover dL (kg) | Zero Error E0 (kg) | Turning Point P (kg) | Raw Error E (kg) | Corrected Error Ec (kg) | Applicable MPE (kg) | Absolute Error ¦Ec¦ (kg) | Decision | OIML Clause | Calculation Formula
1 | LOADING | 0 | 0 | 0.001 | 0.0015 | 0.0015 | 0.0015 | 0 | 0.0025 | 0 | PASS | Clause A.4.4.1 | P = I + 0.5e - dL; E = P - L; Ec = E - E0
2 | LOADING | 3.75 | 3.75 | 0.001 | 0.0015 | 3.7515 | 0.0015 | 0 | 0.005 | 0 | PASS | Clause A.4.4.1 | P = I + 0.5e - dL; E = P - L; Ec = E - E0
3 | LOADING | 7.5 | 7.501 | 0.001 | 0.0015 | 7.5025 | 0.0025 | 0.001 | 0.005 | 0.001 | PASS | Clause A.4.4.1 | P = I + 0.5e - dL; E = P - L; Ec = E - E0
4 | LOADING | 15 | 15.001 | 0.001 | 0.0015 | 15.0025 | 0.0025 | 0.001 | 0.0075 | 0.001 | PASS | Clause A.4.4.1 | P = I + 0.5e - dL; E = P - L; Ec = E - E0

## REPEATABILITY
Series | Reading No | Test Load (kg) | Indication (kg) | Additional dL (kg) | Observation Status
1 | 1 | 7.5 | 7.5 | 0 | RECORDED
1 | 2 | 7.5 | 7.5 | 0 | RECORDED
1 | 3 | 7.5 | 7.5005 | 0 | RECORDED
1 | 4 | 7.5 | 7.5 | 0 | RECORDED
1 | 5 | 7.5 | 7.5 | 0 | RECORDED
1 | 6 | 7.5 | 7.5005 | 0 | RECORDED
1 | 7 | 7.5 | 7.5 | 0 | RECORDED
1 | 8 | 7.5 | 7.5 | 0 | RECORDED
1 | 9 | 7.5 | 7.5005 | 0 | RECORDED
1 | 10 | 7.5 | 7.5 | 0 | RECORDED
REPEATABILITY STATISTICAL SUMMARY | CALCULATED VALUE | APPLICABLE LIMIT | METROLOGICAL DECISION
Maximum Indication (Imax) | 7.5005 kg | - | -
Minimum Indication (Imin) | 7.5000 kg | - | -
Difference dI (Imax - Imin) | 0.0005 kg | 0.0050 kg | PASS
Applicable MPE (¦MPE¦) | 0.0050 kg | Clause 3.6.1 Table 6 | COMPLIANCE BOUND
Difference vs MPE (dI - MPE) | -0.0045 kg | <= 0.0000 kg | PASS
Required Reading Count | 3 readings (or 10 for Class I/II) | Clause A.4.10 | SATISFIED
Actual Reading Count Recorded | 10 readings recorded | >= Required count | SATISFIED
Metrological Determination | Difference dI = 0.0005 kg | MPE = 0.0050 kg | PASS

## ECCENTRIC LOADING
Receptor Position | Applied Load (kg) | Indication (kg) | Changeover dL (kg) | Zero Error E0 (kg) | Raw Error E (kg) | Corrected Error Ec (kg) | MPE Limit (kg) | Decision | OIML Clause
CENTER | 5 | 5 | 0.001 | 0 | 0.0015 | 0.0015 | 0.005 | PASS | Clause 3.6.2 & A.4.7
FRONT_LEFT | 5 | 5.001 | 0.001 | 0 | 0.0025 | 0.0025 | 0.005 | PASS | Clause 3.6.2 & A.4.7
BACK_LEFT | 5 | 5.001 | 0.001 | 0 | 0.0025 | 0.0025 | 0.005 | PASS | Clause 3.6.2 & A.4.7
BACK_RIGHT | 5 | 5.001 | 0.001 | 0 | 0.0025 | 0.0025 | 0.005 | PASS | Clause 3.6.2 & A.4.7
FRONT_RIGHT | 5 | 5.001 | 0.001 | 0 | 0.0025 | 0.0025 | 0.005 | PASS | Clause 3.6.2 & A.4.7

## DISCRIMINATION
Test Point | Base Test Load (kg) | Initial Indication (kg) | Additional Load 1.4d (kg) | Additional Load Formula | Resulting Indication (kg) | Observed Change (kg) | Required Step I + d (kg) | Decision | OIML Clause
Point #1 | 9 | 9 | 0.0014 | Delta_L = 1.4 * d | 9.001 | 0.001 | 9.001 | PASS | Clause 3.8 & A.4.8

## ZERO SETTING
Metrological Parameter | Observation / Calculation | OIML R-76 Limit | Decision | Relevant OIML Clause | Technician Notes
Initial Indication at Zero Load | 0 kg | Nominal 0.000 | RECORDED | Clause 4.5.1 | Zero load applied on platter
Zero-Setting Device Action | Zero-setting key activated | Mandatory device | PASS | Clause 4.5.2 | Zero-setting device successfully brought indicator to zero
Final Indication after Zero-Setting | 0 kg | Zero indication | PASS | Clause 4.5.2 | Digital display reads zero
Calculated Zero Error E0 | 0 kg | +/- 0.00125 kg (+/-0.25e) | PASS | Clause 4.5.2 & A.4.2.1 | E0 = I + 0.5e - dL - L (at zero load)

## OTHER TESTS
Test Code | Test Name | R-76 Reference | Status | Applicability | Verification Summary / Technician Notes
MULTIPLE_INDICATING_DEVICES | Multiple Indicating Devices Comparison | OIML R 76-1:2006 Clause 3.6.3 | NOT_APPLICABLE | NOT_APPLICABLE | Verified
DIFFERENT_POSITIONS_OF_EQUILIBRIUM | Different Positions of Equilibrium | OIML R 76-1:2006 Clause 3.6.4 | NOT_APPLICABLE | NOT_APPLICABLE | Verified

## CALCULATION SUMMARY
Test Module | OIML R-76 Clause | Implementation State | Key Observed Result / Evidence | Applied MPE Limit | Metrological Decision
Weighing Performance & Intrinsic Error | OIML R 76-1:2006 Clause 3.5.1 & A.4.4.3 | IMPLEMENTED | Max error Ec: 0.001 kg | MPE: +/-0.0075 kg | PASS
Repeatability Test | OIML R 76-1:2006 Clause 3.6.1 & A.4.10 | IMPLEMENTED | Delta_I: 0.0005 kg | MPE: 0.005 kg | PASS
Eccentric Loading Test | OIML R 76-1:2006 Clause 3.6.2 & A.4.7 | IMPLEMENTED | Max off-centre error: 0 kg | MPE at 1/3 Max | PASS
Zero-Setting & Zero-Tracking Evaluation | OIML R 76-1:2006 Clause 4.5 & A.4.2 | IMPLEMENTED | Zero error E0: 0 kg | +/-0.25 e | PASS
Discrimination Test | OIML R 76-1:2006 Clause 3.8 & A.4.8 | IMPLEMENTED | 1.4d addition response | I + d step | PASS
Multiple Indicating Devices Comparison | OIML R 76-1:2006 Clause 3.6.3 | IMPLEMENTED | Not applicable to device specification | N/A | NOT_APPLICABLE
Different Positions of Equilibrium | OIML R 76-1:2006 Clause 3.6.4 | IMPLEMENTED | Not applicable to device specification | N/A | NOT_APPLICABLE

## OVERALL COMPLIANCE
COMPLIANCE AUDIT FIELD | AUTHORITATIVE OUTCOME
Session Identifier | N/A
Official Report Number | R76-2026-000073
Public Verification ID | R76-2026-D6F438AF
Regulatory Baseline | OIML R 76-1:2006 (TYPE_EVALUATION)
CANONICAL OVERALL STATUS | PASSED
Compliance Determination | All 5 applicable OIML R 76-1:2006 metrological tests have been verified compliant with Maximum Permissible Errors.
Total Test Modules Evaluated | 7 modules
Compliant Modules (PASS) | 5 modules
Non-Compliant Modules (FAIL) | 0 modules
Review Required Modules | 0 modules
Not Applicable Modules | 2 modules

## AUDIT & TRACEABILITY
AUDIT / TRACEABILITY PARAMETER | RECORDED AUDIT VALUE
Report Generation Timestamp | 2026-09-26T17:05:36.477Z
Testing Officer Name | Pramod Patil
Testing Facility Name | Central Legal Metrology Testing Laboratory
Report Number | R76-2026-000073
Public QR Verification ID | R76-2026-D6F438AF
Public Verification Route URL | http://localhost:5173/public/report/R76-2026-D6F438AF
OIML R-76 Engine Version | Phase 4.1 Strict Canonical Trace Engine
Data Integrity Baseline | Tamper-Evident Canonical Result Flow (Zero Override Allowed)
Generated Artifact Type | Full Observation & Calculation Technical Record (.xlsx)